CTR-125 Startup and Shutdown Procedure

##### Since CTR-125 power is always ON except during system maintenance, these steps are not normally required. System power ON or OFF procedures would normally be conducted by a maintenance engineer. You will not be involved in removing or restoring system power except in an emergency situation. If you are assisting an engineer, under his direction, perform power ON and power OFF according to these instructions.

##### The CTR-125 has two red-colored Emergency OFF buttons (EMO) installed at strategic places on the system. They are located:

- On the front panel

- Behind the system on the back  panel

CTR-125 Startup Process

- Verify that both red EPO (Emergency Power Off) switches (located at either end of the CTR-125) are in the Out position. Twist and pull the EPO switches out to be able to reset power to the system.

- The Reset Key Switch is located on the back end of the unit. Insert and turn the key to the right and hold while turning the Control Power On breaker to the On (up) position. The green Power On indicator should remain illuminated.

- On the front end of the CTR-125, the blue Control Power Standby indicator should be illuminated. Press the green Primary Control On switch to start the system.

- Wait approximately 30 seconds until the Eurotherm indicators on the control panel have initialized and are indicating temperature.

- Press the Over Temp Reset switch.

- Walk around to the back end of the CTR-125 and flip the Process Element On circuit breaker to the on (up) position.

- On the front end of the CTR-125, the blue Heating Element Standby indicator should be illuminated. Press the green Primary Heating Element On switch to start the system.

- The CTR-125 is now ready for operation.

CTR-125 Onsite Startup Process

- Ground Check – With Power feed to unit Off - Test continuity from input ground to other ground points and chassis. Test for open between ground and all power points inside box.

- Turn power to unit On – measure 208V at input cable (barrier block) to neutral and ground.

- Verify that Altech power supply is set to 230V.

- Power-up – Turn key to the right and hold (square green light should come on), then flip Primary Control On breaker up (On).

- Honeywell MFC green light should be on indicating “No Hydrogen.”

- Check Overtemp cutoffs using manual controls of small Eurotherm Overtemp Indicators

a)	Press yellow Overtemp Reset switch – Red temp fault lamp should go out.

b)	Manually lower trip point on Eurotherm overtemp indicators (one at a time) and verify that Overtemp fault trips for each zone. Does this trip the main breaker???

CTR-125 Shutdown Process

- Press the Heating Element On switch to turn it off. The green indicator should go out and the blue Heating Element Standby indicator should be illuminated.

- Press the green Control Power On switch, located on the front end of the unit. The Control Power Standby indicator will illuminate.

- Walk around to the back end of the CTR-125 and turn the Primary Control circuit breaker Off.

- The CTR-125 is now powered down in standby mode.

Caution

- Normal fab procedure requires a nitrogen system purge after a batch has been processed, especially LPCVD units. This prevents any potentially noxious gases from contaminating the area after shutdown.

- Caution: Always confirm that the system has been purged prior to powering down.

Note: Gas On/Off

- The number of gas shutoff valves is dependent on the type of system. The valves are located inside the gas cabinet in the back portion of the system. If the system is powered down for an extended period of time, shut off the gas supplies. The valves are clearly marked with flow on/off direction indicators